## Slide 1: ELPI+™
Electrical Low Pressure Impactor

## Slide 2: ELPI+™ Electrical Low Pressure Impactor
Real-time particle number size distribution and concentration measurement, 10 Hz
6 nm - 10 m, 14 size fractions
Particles are collected – enables subsequent chemical analysis on the collected samples
Dynamic range, from outdoor air to power plant stack concentrations

## Slide 3: ELPI+™ Operating Principle
Operation based on three main
components:
  Impactor
    Particle size fractionation
  Charger
    Particle are charged before fractionating
  Electrometers
    Current distribution - directly proportional to number distribution
    Fast, sensitive

## Slide 4: ELPI+™ Operating Principle
HV

## Slide 5: ELPI+™ Charger
Corona charger
  Unipolar (+)
  Diode-type
Charging
  Corona discharge
  Diffusion charging
5 kV
ELPI+™ charger generates ion plume and particles are charged based on (mainly) diffusion charging.
  Charge of particles is directly related to ion- and condensation sinks and active surface

## Slide 6: Charger Efficiency
Experimental relationship between current and concentration
  Marjamäki et al. 2000
DOS and Diesel soot give same results
  Moisio’s PhD Thesis, 1999
PneQ=fAcm3

## Slide 7: ELPI+™ Impactor
15 stage cascade low-pressure impactor
Small deposit area, 25 mm
40 mbar low pressure
6 nm – 10 μm
Stage 1 (6-16 nm) integrated filter

## Slide 8: Impactor Operating principle
Turn the gas flow sharply
Particles with sufficient inertia cannot follow the flow
Particles with small enough inertia remain in the flow
Capture the particles escaping the flow

## Slide 9: Particle size concept
The diameter of particle varies depending on the measurement method
Real Particle
Stokes diameter
Aerodynamic
diameter
Optical diameter
Bulk density
Unit density
Optical properties
+
Mobility diameter
Charging, electric field
SEM picture
Volatility, coating, “surface”

## Slide 10: Impactor
Aerodynamic diameter
Gas velocity and dimensions different in different stages
Collection plate
Flow
Nozzle Diameter Dj
Jet to plate distance
Jet velocity U

## Slide 11: Collection efficiency - Bouncing
Shape of the distribution changes
 Surface coating, e.g. greasing

## Slide 12: Collection efficiency
0%
10%
20%
30%
40%
50%
60%
70%
80%
90%
100%
0.01
0.1
1.
10.
100.
Dp [µm]
Bounce
Interception,
diffusion
Ideal curve
The sharper the better

## Slide 13: D50 -cutdiameter
D50% (=cut size)
  50% of particles are collected
  corresponding to Stk50%
Dp [µm]
0%
10%
20%
30%
40%
50%
60%
70%
80%
90%
100%
0.01
0.1
1.
10.
100.
50%
D50

## Slide 14: Collection efficiency - overloading
Change in the flow and cutpoints
 Shorter samples or diluting
Collection plate

## Slide 15: Cascade Impactor - example
Stage 3: >10 µm
Stage 1: >1.0 µm
Stage 2: >2.5 µm
Collection substrates
Collection plate
Flow
Nozzle Diameter Dj
Jet to plate distance
Jet velocity U

## Slide 16: Impactor calibration
Good calibration information
  Marjamäki, M., Keskinen, J., Chen, D-R. and Pui, D. Y. H. (2000) Performance Evaluation of the Electrical Low-Pressure Impactor (ELPI), Journal of Aerosol Science 31:2, pp. 249-261
Low particle losses
  Virtanen, A., Marjamäki, M., Ristimäki, J., Keskinen, J. (2001). Fine particle losses in electrical low-pressure impactor, Journal of Aerosol Science, vol. 32, pp. 389-401

## Slide 17: Impactor calibration and publications

## Slide 18: D50 -cutdiameter
Velocity of gas (U)
Size of particle (Dp)
Particle density (ρp)
Size of body (Dj)
Gas viscosity ()
Slip correction factor (Cc)
Dimensionless parameter: Stokes number (Stk) = ratio of particle stopping distance () to collector dimensions

## Slide 19: Impactor calibration
Vacuum pump
PIN
1013,3 mbar
POUT
40 mbar
P15 P14 P13 P12 P11 P10 P9 P8 P7 P6 P5 P4 P3 P2
LFE
  Pressure calibration for each unit
  Cut sizes are calculated from flow rate and pressure drops
  All impactors are calibrated also with particles

## Slide 20: ELPI+™ impactor
14 size fractions
New stage at 16 nm
”Filter stage” standard in all units
See correct values in the data sheet
Stage | D50 | Di
15 | 10 | 
14 | 6.8 | 8.2
13 | 4.4 | 5.5
12 | 2.5 | 3.3
11 | 1.6 | 2.0
10 | 1.0 | 1.3
9 | 0.64 | 0.80
8 | 0.40 | 0.51
7 | 0.26 | 0.32
6 | 0.17 | 0.21
5 | 0.108 | 0.14
4 | 0.060 | 0.080
3 | 0.030 | 0.042
2 | 0.017 | 0.022
1 | 0.006 | 0.010

## Slide 21: ELPI+™
Good calibration information
  Marjamäki, M., Keskinen, J., Chen, D-R. and Pui, D. Y. H. (2000) Performance Evaluation of the Electrical Low-Pressure Impactor (ELPI), Journal of Aerosol Science 31:2, pp. 249-261.
  Pressure calibration for each unit
Low losses
  Virtanen, A., Marjamäki, M., Ristimäki, J., Keskinen, J. (2001). Fine particle losses in electrical low-pressure impactor, Journal of Aerosol Science, vol 32, pp 389-401.

## Slide 22: Impactor stage
Current measurement from outside of the stage
Whole Impactor stage acts as Faradays cage
Collection material does not matter
Q
cable
Electrometer
IIN -IOUT
IOUT
IIN

## Slide 23: Stage construction
Jet plate
PEEKTM Insulator with O-rings
Collection plate

## Slide 24: ELPI+™ samples from Santiago (city center)
Collection time: ~4 days

## Slide 25: ELPI+™ Operating Principle
HV

## Slide 26: Electrometers
Current measurement
Gas-sealed electrometer case
0.1 to 500 000 fA (10-15 A)
One measurement range

## Slide 27: ELPI+™ stand alone unit
ELPI+™ display helps you:
  See online data
  Control the unit
  Measure
  Save the data
ELPI+VI software for Laptops
  ELPI+™ can be used with laptop
  Controlling ELPI+™
  Save the data
  Automatic Q/N and Q/M measurement
  Measure

## Slide 28: ELPI+™ calculation
Inside ELPI+™ and ELPI+VI
Data file format similar to ELPI™
Some additional values
  E.g. Charger voltage, charger current, impactor low pressure, 6 AUX channel values, sample air temperature and pressure
  Raw values saved in the data file as before
Data processing in ELPIPlusCalc.xls
Description of the method in the ELPI+™ user manual

## Slide 29: ELPI+™ Calculation
ELPI+™ measures currents carried by particles and impacting to impactor stages
ELPI+™ calculates particles
  Number concentration
  Mass concentration
  Volume concentration
  Etc.

## Slide 30: ELPI+™ Calculation
(Marjamäki et. al, Journal of Aerosol Sci. v30 n. 1 2000)
PneQ=fAcm3

## Slide 31: ELPI+™ Calculation
From number size concentration we know
  Number size concentration
  Particle size
  How many particles
From number size concentration can be calculated
  Surface concentration
  Volume concentration
  Mass concentration

## Slide 32: ELPI+™ Display
Instrument status
Size distribution
Concentration
Menu controls

## Slide 33: ELPI+™ Display

## Slide 34: ELPI+™ operation via ELPI+VI
ELPI+VI for data display and instrument control
  Data collected in real-time
  Can be analysed with ELPIPlusCalc.xls
Simple user interface
Automatic Q/N and Q/M measurement

## Slide 35: ELPI+™ applications
Combustion research
Outdoor air quality measurements
Indoor air
Automotive testing and research
Blow-by gas
Brake wear debris
Pharmaceutical studies
Welding fumes
…
Wide variety of applications
Application notes/papers available
Publications list at www.dekati.com

## Slide 36: ELPI+™
Real time measurement of
  Number concentration
  Mass concentration
  Active surface area
  Ion sink
  Condensation sink
  → Describes activity of particle
Particle collection
  Chemical and Physical analysis
How particles are growing or changing?
How harmful particles can be?
What kind of particles?

## Slide 37: Collection: Mass and Chemical analysis
© Dekati Ltd.
Mass and chemical analysis of size classified particles
D. Temesi et al: Size resolved chemical mass balance of aerosol particles over rural Hungary. Atmospheric Environment 35 (2001) 4347–4355

## Slide 38: Collection: SEM and TEM
Karsten Fuglsang, Jacob Markussen, Karina Brotoft, Thue G. Frederiksen
FORCE Technology
ETH Conference on Combustion Generated Nanoparticles, Zürich, 4 Aug. 2010

## Slide 39: Active surface of particles
The part of surface of particle which can be in contact with lung cells is called active surface
Health effects correlate with active surface of particles
Active surface is related to processes related to surface
  Condensing (growing) – Condensation sink
  Charging – Ion sink
Can be calculated using mobility size based on theory
Can be measured directly using diffusion charger and current measurement

## Slide 40: ELPI™ Charger
Corona charger
  Unipolar (+)
  Diode-type
Charging
  Corona discharge
  Diffusion charging
5 kV
ELPI+ charger generates ion plume and particles are charged based on (mainly) diffusion charging.
  Charge of particles is directly related to ion- and condensation sinks and active surface

## Slide 41: Active surface, Ion - and condensation sink
Active surface
  Atot=constant x ELPI current
  Willeke & Baron, Aerosol Measurement, 2nd edition
ELPI+ total current is related to ion and condensation sink.
  Ion sink: = 8.55E-06 s−1fA−1 x ELPI+ total current
  Condensation sink = 7.27E-06 s−1fA−1 x ELPI+ total current
  Kuuluvainen, H., Kannosto, J., Virtanen, A., Mäkelä, J. M., Kulmala, M., Aalto, P., Keskinen, J. , Technical Note: Measuring condensation sink and ion sink of atmospheric aerosols with the electrical low pressure impactor (ELPI), Atmos. Chem. Phys., 10, 1361-1368, 2010.

## Slide 42: Gravimetric measurements with the ELPI+™

## Slide 43: Gravimetric measurements
ELPI+™ Impactor can be used as a gravimetric impactor
Aluminium/polycarbonate foils
Charger off
Weigh the substrates

## Slide 44: Particle charge measurements

## Slide 45: ELPI+™ for charge distribution measurements
Achieved by turning ELPI+™ charger off
  Readings in fA (10-15 A)
  1A = 1 Coulomb /s
Charge distribution of aerosol in real time
  Net charge of particles
For charge/number (Q/N) result charger can be switched on and off (~5-10 seconds) or parallel measurements can be conducted
 Automatic Q/N and Q/M calculation in ELPI+VI
  Automatic charger ON and OFF measurement

## Slide 46: Particle density measurements

## Slide 47: Aerodynamic size vs. Mobility size
da
Aerodynamic size
dme
Mobility size
Density
1 g/cm3
Density
2 g/cm3
Density
0.85 g/cm3
Particle size is difficult to measure directly. Measurement instruments measure a quantity which is related to particle size.
Density increases, da increases.
Density decreases, da decreases.
1 g/cm3
Different instruments measure different quantity → related to different particle size.

## Slide 48
Density
1 g/cm3
Density
2 g/cm3
Density
0.85 g/cm3
Effective density takes account shape of particle and cavity
Different materials can have different density
The effective density of non-spherical particle is alway smaller than bulk/material density of particle
Density of particle can give information about:
material
shape
Growth and formation processes of particle
Effective density
< 1 g/cm3
Effective density of particle

## Slide 49: ELPI+™
Real time measurement of
  Number concentration
  Mass concentration
  Active surface area
  Ion sink
  Condensation sink
  → Describes activity of particle
Particle collection
  Chemical and Physical analysis
How particles are growing or changing?
How harmful particles can be?
What kind of particles?

## Slide 50: Air Quality
Outdoor – Indoor – Workplace

## Slide 51
Air Quality
Visibility
Climate change
Atmospheric research
Weather forecast
Health effects

## Slide 52: Air Quality
Cooking
Bacteria
Viruses
Emissions
Mold
Smoke
Asthma / Allergy Trigger

## Slide 53
Air Quality: Indoor and workplace
Exposure
Legislation
Particle limits for:
Number?
Mass?
Surface?
Composition?
Safety issues
Health effects
Dose

## Slide 54: Nanotechnology
It is everywhere
But…
Are nanoparticles harmful and how harmful?
What are the main sources and emissions?
What are health relevant properties?
  Number? Mass? Active surface area?
What needs to be known:
emission-concentration-exposure-dose-effect

## Slide 55: Air quality: Health effects
premature death of people with heart or lung disease
nonfatal heart attacks
irregular heartbeat
asthma
decreased lung function
increased respiratory symptoms, such as irritation of the airways, coughing or difficulty with breathing
premature death / year
350 000 in Europe
22 000-52 000 in USA
1300 in Finland
The scientific evidence that air pollution causes disease is no longer in doubt, according to Dr Michal Krzyzanowski, an epidemiologist working until recently at World Health Organization. He says:“Circulatory and respiratory diseases associated with exposure to air pollution lead to a reduction in life expectancy of 10 months in the Polish population.

## Slide 56: Dekati® Solutions
Engine Emissions
After-treatment device and engine development
On-board measurements
Blow-by gas emissions
EURO5b regulatory measurements

## Slide 57: Dekati® Solutions
Emissions measurements
  After-treatment device development
  Combustion optimization and research
  Real-time monitoring of particle size
  ISO23210 measurements of PM10 and PM2.5